## Sheet1
Opto-board Assembly Status
0
Opto-board S/N: |  |  |  | Site(check one): |  | OSU |  | Seigen
Opto-board flavor (check one): |  |  | B-Layer |  | Left |  | Right
VDC S/N | 0
0
DORIC S/N |  |  | -- | --
Step |  |  |  | Date | Initial | Comments
1) Reinforce 80-pin Connector
2) Mechanical tolerance test
3) short test
4) VDC mounting
5) DORIC mounting
6) Measure VCSEL power
7) Measure PIN respnsivity
8) Glue VCSEL opto-pack (?)
9) Glue Housing
10.1) Glue 2nd VCSEL opto-pack(?)
10.2) Glue housing (?)
11.1) Glue PIN opto-pack
11.2) Glue housing
12) Solder VCSEL opto-pack
13) Solder 2nd VCSEL opto-pack (?)
14) Solder PIN opto-pack
15) Wire bonding
16) Basic Functionality
17) Encapsulation

## Sheet2
Opto-Board QA data Sheet: Optical Power and Responsivity
Opto-board S/N
0
VCSEL Opto-pack S/N:
Power (uW) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Date | Initial | P> 500 uW & dP ,20% ?
Pre Opto-pack Burn-in |  | 846.3573 | 860.2059 | 825.1537 | 938.1145 | 860.9852 | 893.5283 | 913.7409 | -- | -- | --
Post Opto-pack Burn-in | [8 empty] | -- | -- | --
Change (%) | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -- | -- | --
Pre Opto-board Burn-in |  | 1143.9846 | 960.1952 | 1095.3056 | 1138.6183 | 961.3259 | 1042.8352 | 901.4037 | -- | -- | --
Change (%) |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -- | -- | --
Post Opto-board Burn-in |  | 1128.7702 | 981.2921 | 1153.4408 | 1162.3577 | 982.2575 | 1066.4509 | 900.3546 | -- | -- | --
Change (%) |  | -1.3299 | 2.1971 | 5.3077 | 2.0849 | 2.1774 | 2.2646 | -0.11638 | -- | -- | --
Opto-board at 0C |  | 1403.7148 | 1132.4484 | 1414.7543 | 1430.8454 | 1302.0059 | 1195.6598 | 1198.802 | -- | -- | --
Change (%) |  | 22.704 | 17.9394 | 29.1653 | 25.6651 | 35.4386 | 14.6547 | 32.9928 | -- | -- | --
Opto-board at 50C |  | 1169.6659 | 951.0048 | 1174.9417 | 1195.9037 | 1069.2505 | 1004.286 | 1021.0368 | -- | -- | --
Change (%) |  | 2.2449 | -0.95714 | 7.2707 | 5.0311 | 11.2266 | -3.6966 | 13.2719 | -- | -- | --
Opto-board QA (10C) |  | 1362.1584 | 1093.9552 | 1366.3368 | 1386.0377 | 1249.3369 | 1170.5463 | 1185.0451 | -- | -- | --
Change (%) |  | 19.0714 | 13.9305 | 24.7448 | 21.7298 | 29.9598 | 12.2465 | 31.4666 | -- | -- | --
VCSEL Opto-pack S/N:
|  | 947.6881 | 873.0093 | 918.7188 | 991.4432 | 915.9322 | 957.2437 | 950.1784
Power (uW) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Date | Initial | P> 500 uW & dP ,20% ?
Pre Opto-pack Burn-in | [8 empty] | -- | -- | --
Post Opto-pack Burn-in | [8 empty] | -- | -- | --
Change (%) | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -- | -- | --
Pre Opto-board Burn-in |  | 1261.2497 | 1178.3895 | 1319.9834 | 1315.2739 | 1402.3166 | 1244.2657 | 1220.9406 | -- | -- | --
Change (%) |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -- | -- | --
Post Opto-board Burn-in |  | 1335.8751 | 1063.2193 | 1303.1479 | 1337.2895 | 1398.8545 | 1171.9405 | 1190.6322 | -- | -- | --
Change (%) |  | 5.9168 | -9.7735 | -1.2754 | 1.6738 | -0.24688 | -5.8127 | -2.4824 | -- | -- | --
Opto-board at 0C |  | 1408.6445 | 1291.8976 | 1468.5029 | 1465.0018 | 1647.2567 | 1489.6986 | 1537.7658 | -- | -- | --
Change (%) |  | 11.6864 | 9.6325 | 11.2516 | 11.3838 | 17.4668 | 19.7251 | 25.9493 | -- | -- | --
Opto-board at 50C |  | 1239.0226 | 1119.9511 | 1289.6229 | 1267.7893 | 1417.1397 | 1261.7108 | 1290.6178 | -- | -- | --
Change (%) |  | -1.7623 | -4.9592 | -2.3001 | -3.6102 | 1.057 | 1.402 | 5.7068 | -- | -- | --
Opto-board QA (10C) |  | 1410.1778 | 1249.3377 | 1453.3409 | 1426.1049 | 1596.0024 | 1437.9025 | 1505.0934 | -- | -- | --
Change (%) |  | 11.808 | 6.0208 | 10.103 | 8.4265 | 13.8118 | 15.5623 | 23.2733 | -- | -- | --

## Sheet3
PIN Opto-pack S/N:
Responsivity (A/W) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Date | Initial | R> 0.5 A/W & dR <5%
Opto-pack QA | [8 empty] | -- | -- | --
Pre-mount | [8 empty] | -- | -- | --
Change (%)

## Sheet4
Opto-board QA Data Sheet
Opto-board S/N:
0
Current consumption:
| Unit | Min | Max | I |  |  | dI <5% | Date |  | Initial |  | Pass/Fail (?)
Pre-QA | mA | 342 | 378 | 354.4922 |  |  |  | 2020-06-18 |  |  |  | Pass
QA | mA | 342 | 378 | 348.7549 |  |  |  | 2020-06-02 |  |  |  | Pass
*The allowable range for current consumption for a B-layer opto-board is 330-400 mA.
Dark current:
| Unit | Max | I |  | Date |  | Initial |  | Pass/Fail (?)
Pre-QA | uA | 1 | 0.36544 |  | 2020-06-18 |  |  |  | Pass
QA | uA | 1 | 0.13614 |  | 2020-06-02 |  |  |  | Pass
Optical
| Unit | Max | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | Pass/Fail (?)
Rise | ns | 1 | 0.37764 | 0.32159 | 0.46729 | 0.32852 | 0.55038 | 0.39714 | 0.36993 | 0.37749 | 0.48574 | 0.68968 | 0.58383 | 0.51093 | 0.46447 | 0.39916 |  |  | Pass
Fall | ns | 1 | 0.59929 | 0.65788 | 0.73672 | 0.86611 | 0.83012 | 0.90304 | 0.56824 | 0.59652 | 0.73218 | 0.80746 | 0.56978 | 0.62333 | 0.44542 | 0.47428 |  |  | Pass

## optical calc
Post Mounting @ room Temp
Power
Power 2
Power 3
channel# | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
avg | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
-10C
Power
Power 2
Power 3
channel# | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
avg | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
10C
Power
Power 2
Power 3
channel# | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
avg | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
50C
Power
Power 2
Power 3
channel# | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
avg | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000

## Sheet5
Electrical: |  |  |  | Date (DD/MM/YY):06/02/2020 | [5 empty] | Initial:
| Unit | Min | Max | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | Pass/fail(?)
Clk Rise | ns | -- | 1 | Date (DD/MM/YY):06/02/2020 1: 0.62117; Date (DD/MM/YY):06/02/2020 2: 0.62474; Date (DD/MM/YY):06/02/2020 3: 0.64396; Date (DD/MM/YY):06/02/2020 4: 0.65194; Date (DD/MM/YY):06/02/2020 5: 0.71015; Date (DD/MM/YY):06/02/2020 6: 0.71078; Initial: 7: 0.68393 | Pass
Clk Fall | ns | -- | 1 | Date (DD/MM/YY):06/02/2020 1: 0.61188; Date (DD/MM/YY):06/02/2020 2: 0.61788; Date (DD/MM/YY):06/02/2020 3: 0.66658; Date (DD/MM/YY):06/02/2020 4: 0.67418; Date (DD/MM/YY):06/02/2020 5: 0.68311; Date (DD/MM/YY):06/02/2020 6: 0.68549; Initial: 7: 0.70327 | Pass
Data Rise | ns | -- | 1 | Date (DD/MM/YY):06/02/2020 1: 0.59351; Date (DD/MM/YY):06/02/2020 2: 0.59863; Date (DD/MM/YY):06/02/2020 3: 0.5984; Date (DD/MM/YY):06/02/2020 4: 0.60873; Date (DD/MM/YY):06/02/2020 5: 0.62334; Date (DD/MM/YY):06/02/2020 6: 0.62675; Initial: 7: 0.64398 | Pass
Data Fall | ns | -- | 1 | Date (DD/MM/YY):06/02/2020 1: 0.68873; Date (DD/MM/YY):06/02/2020 2: 0.68609; Date (DD/MM/YY):06/02/2020 3: 0.69216; Date (DD/MM/YY):06/02/2020 4: 0.71382; Date (DD/MM/YY):06/02/2020 5: 0.71366; Date (DD/MM/YY):06/02/2020 6: 0.7316; Initial: 7: 0.77176 | Pass
Jitter @ IPIN = 100uA | ns | -- | 1 | Date (DD/MM/YY):06/02/2020 1: 0.20163; Date (DD/MM/YY):06/02/2020 2: 0.07555; Date (DD/MM/YY):06/02/2020 3: 0.05454; Date (DD/MM/YY):06/02/2020 4: 0.07578; Date (DD/MM/YY):06/02/2020 5: 0.09241; Date (DD/MM/YY):06/02/2020 6: 0.17644; Initial: 7: 0.17229 | Pass
Jitter @ IPIN = 1000uA | ns | -- | 1 | Date (DD/MM/YY):06/02/2020 1: 0.19269; Date (DD/MM/YY):06/02/2020 2: 0.06097; Date (DD/MM/YY):06/02/2020 3: 0.05649; Date (DD/MM/YY):06/02/2020 4: 0.06513; Date (DD/MM/YY):06/02/2020 5: 0.08492; Date (DD/MM/YY):06/02/2020 6: 0.17758; Initial: 7: 0.16105 | Pass
Pos Width max @ IPIN = 100uA |  |  |  | Date (DD/MM/YY):06/02/2020 1: 12.666; Date (DD/MM/YY):06/02/2020 2: 12.965; Date (DD/MM/YY):06/02/2020 3: 12.468; Date (DD/MM/YY):06/02/2020 4: 12.153; Date (DD/MM/YY):06/02/2020 5: 13.053; Date (DD/MM/YY):06/02/2020 6: 12.472; Initial: 7: 13.228
Pos Width max @ IPIN = 1000uA |  |  |  | Date (DD/MM/YY):06/02/2020 1: 12.591; Date (DD/MM/YY):06/02/2020 2: 12.844; Date (DD/MM/YY):06/02/2020 3: 12.499; Date (DD/MM/YY):06/02/2020 4: 12.146; Date (DD/MM/YY):06/02/2020 5: 12.976; Date (DD/MM/YY):06/02/2020 6: 12.313; Initial: 7: 13.109
Pos Width min @ IPIN = 100uA |  |  |  | Date (DD/MM/YY):06/02/2020 1: 11.873; Date (DD/MM/YY):06/02/2020 2: 12.46; Date (DD/MM/YY):06/02/2020 3: 12.096; Date (DD/MM/YY):06/02/2020 4: 11.705; Date (DD/MM/YY):06/02/2020 5: 12.511; Date (DD/MM/YY):06/02/2020 6: 11.709; Initial: 7: 12.468
Pos Width min @ IPIN = 1000uA |  |  |  | Date (DD/MM/YY):06/02/2020 1: 11.931; Date (DD/MM/YY):06/02/2020 2: 12.497; Date (DD/MM/YY):06/02/2020 3: 12.127; Date (DD/MM/YY):06/02/2020 4: 11.768; Date (DD/MM/YY):06/02/2020 5: 12.559; Date (DD/MM/YY):06/02/2020 6: 11.721; Initial: 7: 12.515
Pos Width mean @ IPIN = 100uA |  |  |  | Date (DD/MM/YY):06/02/2020 1: 12.2618; Date (DD/MM/YY):06/02/2020 2: 12.6764; Date (DD/MM/YY):06/02/2020 3: 12.2918; Date (DD/MM/YY):06/02/2020 4: 11.9412; Date (DD/MM/YY):06/02/2020 5: 12.7584; Date (DD/MM/YY):06/02/2020 6: 12.0376; Initial: 7: 12.8299
Pos Width mmean @ IPIN = 1000uA |  |  |  | Date (DD/MM/YY):06/02/2020 1: 12.254; Date (DD/MM/YY):06/02/2020 2: 12.6518; Date (DD/MM/YY):06/02/2020 3: 12.2875; Date (DD/MM/YY):06/02/2020 4: 11.9313; Date (DD/MM/YY):06/02/2020 5: 12.7632; Date (DD/MM/YY):06/02/2020 6: 12.0206; Initial: 7: 12.8247
GoNoGoDuty Cycle | % | 46 | 54 | Date (DD/MM/YY):06/02/2020 1: 49.0294; Date (DD/MM/YY):06/02/2020 2: 50.761; Date (DD/MM/YY):06/02/2020 3: 49.1326; Date (DD/MM/YY):06/02/2020 4: 47.7014; Date (DD/MM/YY):06/02/2020 5: 51.062; Date (DD/MM/YY):06/02/2020 6: 48.0889; Initial: 7: 51.2566 | Pass
Duty Cycle @ IPIN = 50uA | % | 46 | 54 | Date (DD/MM/YY):06/02/2020 1: 49.0889; Date (DD/MM/YY):06/02/2020 2: 50.7187; Date (DD/MM/YY):06/02/2020 3: 49.1956; Date (DD/MM/YY):06/02/2020 4: 47.769; Date (DD/MM/YY):06/02/2020 5: 51.0847; Date (DD/MM/YY):06/02/2020 6: 48.1949; Initial: 7: 51.3815 | Pass
Duty Cycle @ IPIN = 1000uA | % | 46 | 54 | Date (DD/MM/YY):06/02/2020 1: 49.0325; Date (DD/MM/YY):06/02/2020 2: 50.6051; Date (DD/MM/YY):06/02/2020 3: 49.1558; Date (DD/MM/YY):06/02/2020 4: 47.7395; Date (DD/MM/YY):06/02/2020 5: 51.0602; Date (DD/MM/YY):06/02/2020 6: 48.0847; Initial: 7: 51.2985 | Pass
Clock+ common mode | V | 1 | 1.5 | Date (DD/MM/YY):06/02/2020 1: 1.1653; Date (DD/MM/YY):06/02/2020 2: 1.1692; Date (DD/MM/YY):06/02/2020 3: 1.1714; Date (DD/MM/YY):06/02/2020 4: 1.1781; Date (DD/MM/YY):06/02/2020 5: 1.192; Date (DD/MM/YY):06/02/2020 6: 1.1656; Initial: 7: 1.1669 | Pass
Clock- common mode | V | 1 | 1.5 | Date (DD/MM/YY):06/02/2020 1: 1.1697; Date (DD/MM/YY):06/02/2020 2: 1.1602; Date (DD/MM/YY):06/02/2020 3: 1.1751; Date (DD/MM/YY):06/02/2020 4: 1.1917; Date (DD/MM/YY):06/02/2020 5: 1.1826; Date (DD/MM/YY):06/02/2020 6: 1.1752; Initial: 7: 1.1534 | Pass
Data+ common mode | V | 1 | 1.5 | Date (DD/MM/YY):06/02/2020 1: 1.0882; Date (DD/MM/YY):06/02/2020 2: 1.0866; Date (DD/MM/YY):06/02/2020 3: 1.0936; Date (DD/MM/YY):06/02/2020 4: 1.1027; Date (DD/MM/YY):06/02/2020 5: 1.0872; Date (DD/MM/YY):06/02/2020 6: 1.0936; Initial: 7: 1.0937 | Pass
Data- common mode | V | 1 | 1.5 | Date (DD/MM/YY):06/02/2020 1: 1.0625; Date (DD/MM/YY):06/02/2020 2: 1.0535; Date (DD/MM/YY):06/02/2020 3: 1.0579; Date (DD/MM/YY):06/02/2020 4: 1.0617; Date (DD/MM/YY):06/02/2020 5: 1.0623; Date (DD/MM/YY):06/02/2020 6: 1.0569; Initial: 7: 1.0611 | Pass
Clock+ amplitude | V | 0.2 | 0.5 | Date (DD/MM/YY):06/02/2020 1: 0.44385; Date (DD/MM/YY):06/02/2020 2: 0.44293; Date (DD/MM/YY):06/02/2020 3: 0.44357; Date (DD/MM/YY):06/02/2020 4: 0.44727; Date (DD/MM/YY):06/02/2020 5: 0.44417; Date (DD/MM/YY):06/02/2020 6: 0.44132; Initial: 7: 0.43807 | Pass
Clock- amplitude | V | 0.2 | 0.5 | Date (DD/MM/YY):06/02/2020 1: 0.43034; Date (DD/MM/YY):06/02/2020 2: 0.42844; Date (DD/MM/YY):06/02/2020 3: 0.43304; Date (DD/MM/YY):06/02/2020 4: 0.43404; Date (DD/MM/YY):06/02/2020 5: 0.44192; Date (DD/MM/YY):06/02/2020 6: 0.43576; Initial: 7: 0.43407 | Pass
Data+ amplitude | V | 0.2 | 0.5 | Date (DD/MM/YY):06/02/2020 1: 0.36877; Date (DD/MM/YY):06/02/2020 2: 0.36541; Date (DD/MM/YY):06/02/2020 3: 0.36599; Date (DD/MM/YY):06/02/2020 4: 0.36678; Date (DD/MM/YY):06/02/2020 5: 0.36356; Date (DD/MM/YY):06/02/2020 6: 0.36558; Initial: 7: 0.36601 | Pass
Data- amplitude | V | 0.2 | 0.5 | Date (DD/MM/YY):06/02/2020 1: 0.39668; Date (DD/MM/YY):06/02/2020 2: 0.39219; Date (DD/MM/YY):06/02/2020 3: 0.39683; Date (DD/MM/YY):06/02/2020 4: 0.39903; Date (DD/MM/YY):06/02/2020 5: 0.39403; Date (DD/MM/YY):06/02/2020 6: 0.39482; Initial: 7: 0.39675 | Pass
Bit errors @ IPIN = 50uA | -- | -- | 0 | Date (DD/MM/YY):06/02/2020 1: 0; Date (DD/MM/YY):06/02/2020 2: 0; Date (DD/MM/YY):06/02/2020 3: 0; Date (DD/MM/YY):06/02/2020 4: 0; Date (DD/MM/YY):06/02/2020 5: 0; Date (DD/MM/YY):06/02/2020 6: 0; Initial: 7: 0 | Pass
Clk errors @ IPIN = 50uA | -- | -- | 0 | Date (DD/MM/YY):06/02/2020 1: 0; Date (DD/MM/YY):06/02/2020 2: 0; Date (DD/MM/YY):06/02/2020 3: 0; Date (DD/MM/YY):06/02/2020 4: 0; Date (DD/MM/YY):06/02/2020 5: 0; Date (DD/MM/YY):06/02/2020 6: 0; Initial: 7: 0 | Pass
Bit errors @ IPIN = 1000uA | -- | -- | 0 | Date (DD/MM/YY):06/02/2020 1: 0; Date (DD/MM/YY):06/02/2020 2: 0; Date (DD/MM/YY):06/02/2020 3: 0; Date (DD/MM/YY):06/02/2020 4: 0; Date (DD/MM/YY):06/02/2020 5: 0; Date (DD/MM/YY):06/02/2020 6: 0; Initial: 7: 0 | Pass
Clk errors @ IPIN = 1000uA | -- | -- | 0 | Date (DD/MM/YY):06/02/2020 1: 0; Date (DD/MM/YY):06/02/2020 2: 0; Date (DD/MM/YY):06/02/2020 3: 0; Date (DD/MM/YY):06/02/2020 4: 0; Date (DD/MM/YY):06/02/2020 5: 0; Date (DD/MM/YY):06/02/2020 6: 0; Initial: 7: 0 | Pass
VDC Bit errors | -- | -- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Pass