## Sheet1
Opto-board Assembly Status
0
Opto-board S/N: |  |  |  | Site(check one): |  | OSU |  | Seigen
Opto-board flavor (check one): |  |  | B-Layer |  | Left |  | Right
VDC S/N | 0
0
DORIC S/N |  |  | -- | --
Step |  |  |  | Date | Initial | Comments
1) Reinforce 80-pin Connector
2) Mechanical tolerance test
3) short test
4) VDC mounting
5) DORIC mounting
6) Measure VCSEL power
7) Measure PIN respnsivity
8) Glue VCSEL opto-pack (?)
9) Glue Housing
10.1) Glue 2nd VCSEL opto-pack(?)
10.2) Glue housing (?)
11.1) Glue PIN opto-pack
11.2) Glue housing
12) Solder VCSEL opto-pack
13) Solder 2nd VCSEL opto-pack (?)
14) Solder PIN opto-pack
15) Wire bonding
16) Basic Functionality
17) Encapsulation

## Sheet2
Opto-Board QA data Sheet: Optical Power and Responsivity
Opto-board S/N
0
VCSEL Opto-pack S/N:
Power (uW) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Date | Initial | P> 500 uW & dP ,20% ?
Pre Opto-pack Burn-in |  | 1841.1044 | 1779.4047 | 1829.0332 | 1840.657 | 1878.9981 | 1903.0521 | 1765.6638 | -- | -- | --
Post Opto-pack Burn-in |  | 1921.0918 | 1846.1687 | 1876.3093 | 1840.2753 | 1898.4594 | 1938.19 | 1797.0877 | -- | -- | --
Change (%) | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -- | -- | --
Pre Opto-board Burn-in |  | 1596.6165 | 1197.7474 | 1534.3927 | 1442.2175 | 1447.1016 | 1228.7583 | 1421.6103 | -- | -- | --
Change (%) |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -- | -- | --
Post Opto-board Burn-in |  | 1426.8057 | 1087.77 | 1416.446 | 1320.1817 | 1309.7106 | 1332.5023 | 1448.2787 | -- | -- | --
Change (%) |  | -10.6357 | -9.182 | -7.6869 | -8.4617 | -9.4942 | 8.443 | 1.8759 | -- | -- | --
Opto-board at 0C |  | 1428.7401 | 1229.0211 | 1711.2726 | 1462.8675 | 1599.4411 | 1427.9079 | 1547.7316 | -- | -- | --
Change (%) |  | -10.5145 | 2.611 | 11.5277 | 1.4318 | 10.5272 | 16.2074 | 8.8717 | -- | -- | --
Opto-board at 50C |  | 1291.0693 | 1090.383 | 1520.3942 | 1281.2658 | 1434.001 | 1268.1313 | 1363.5346 | -- | -- | --
Change (%) |  | -19.1372 | -8.9639 | -0.91231 | -11.16 | -0.90529 | 3.2043 | -4.0852 | -- | -- | --
Opto-board QA (10C) |  | 1439.5243 | 1191.9196 | 1629.1819 | 1459.5693 | 1575.6202 | 1421.1333 | 1544.7775 | -- | -- | --
Change (%) |  | -9.8391 | -0.48656 | 6.1776 | 1.2031 | 8.8811 | 15.6561 | 8.6639 | -- | -- | --
Reception Test |  | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | -- | -- | --
VCSEL Opto-pack S/N:
Power (uW) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Date | Initial | P> 500 uW & dP ,20% ?
Pre Opto-pack Burn-in |  | 1783.3978 | 1642.565 | 1691.8559 | 1656.9463 | 1671.5156 | 1781.4923 | 1657.866 | -- | -- | --
Post Opto-pack Burn-in |  | 1884.3225 | 1714.4501 | 1732.6827 | 1708.2379 | 1722.3458 | 1750.0232 | 1643.3581 | -- | -- | --
Change (%) | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -- | -- | --
Pre Opto-board Burn-in |  | 1168.7699 | 1019.5603 | 1039.157 | 1066.4896 | 1025.854 | 1140.6607 | 1055.582 | -- | -- | --
Change (%) |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -- | -- | --
Post Opto-board Burn-in |  | 1112.5205 | 1031.793 | 1099.6363 | 1041.4799 | 988.4324 | 1084.5294 | 1097.2112 | -- | -- | --
Change (%) |  | -4.8127 | 1.1998 | 5.82 | -2.3451 | -3.6478 | -4.9209 | 3.9437 | -- | -- | --
Opto-board at 0C |  | 1199.1253 | 1059.0643 | 1241.426 | 1096.4132 | 1012.3346 | 1112.1444 | 1215.1765 | -- | -- | --
Change (%) |  | 2.5972 | 3.8746 | 19.4647 | 2.8058 | -1.3179 | -2.5 | 15.1191 | -- | -- | --
Opto-board at 50C |  | 1000.6075 | 914.0852 | 1048.0622 | 945.4545 | 856.0541 | 984.8771 | 1061.9582 | -- | -- | --
Change (%) |  | -14.388 | -10.3452 | 0.85696 | -11.3489 | -16.5521 | -13.6573 | 0.60405 | -- | -- | --
Opto-board QA (10C) |  | 1213.3927 | 1063.9073 | 1101.4377 | 1047.1652 | 986.2784 | 1078.0412 | 1201.4634 | -- | -- | --
Change (%) |  | 3.8179 | 4.3496 | 5.9934 | -1.812 | -3.8578 | -5.4898 | 13.82 | -- | -- | --
Reception Test |  | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | -- | -- | --

## Sheet3
PIN Opto-pack S/N:
Responsivity (A/W) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Date | Initial | R> 0.5 A/W & dR <5%
Opto-pack QA | [8 empty] | -- | -- | --
Pre-mount | [8 empty] | -- | -- | --
Change (%)

## Sheet4
Opto-board QA Data Sheet
Opto-board S/N:
0
Current consumption:
| Unit | Min | Max | I |  |  | dI <5% | Date |  | Initial |  | Pass/Fail (?)
Pre-QA | mA | 322 | 384 | 367.7979 |  |  |  | 2021-02-22 |  |  |  | Pass
QA | mA | 304 | 366 | 340.7471 |  |  |  | 2021-02-22 |  |  |  | Pass
*The allowable range for current consumption for a B-layer opto-board is 330-400 mA.
Dark current:
| Unit | Max | I |  | Date |  | Initial |  | Pass/Fail (?)
Pre-QA | uA | 1 | 0.56191 |  | 2021-02-22 |  |  |  | Pass
QA | uA | 1 | 0.58252 |  | 2021-02-22 |  |  |  | Pass
Optical
| Unit | Max | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | Pass/Fail (?)
Rise | ns | 1 | 0.33748 | 0.30731 | 0.49678 | 0.43828 | 0.6511 | 0.5074 | 0.26909 | 0.52495 | 0.57997 | 0.50332 | 0.35562 | 0.54948 | 0.44071 | 0.45958 |  |  | Pass
Fall | ns | 1 | 0.49697 | 0.6634 | 0.46306 | 0.41659 | 0.59605 | 0.49753 | 0.68906 | 0.64427 | 0.4669 | 0.46732 | 0.49842 | 0.51571 | 0.46258 | 0.48009 |  |  | Pass

## optical calc
Post Mounting @ room Temp
Power
Power 2
Power 3
channel# | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
avg | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
-10C
Power
Power 2
Power 3
channel# | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
avg | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
10C
Power
Power 2
Power 3
channel# | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
avg | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
50C
Power
Power 2
Power 3
channel# | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
avg | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000

## Sheet5
Electrical: |  |  |  | Date (DD/MM/YY):02/22/2021 | [5 empty] | Initial:
| Unit | Min | Max | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | Pass/fail(?)
Clk Rise | ns | -- | 1 | Date (DD/MM/YY):02/22/2021 1: 0.47885; Date (DD/MM/YY):02/22/2021 2: 0.49546; Date (DD/MM/YY):02/22/2021 3: 0.50643; Date (DD/MM/YY):02/22/2021 4: 0.52912; Date (DD/MM/YY):02/22/2021 5: 0.53274; Date (DD/MM/YY):02/22/2021 6: 0.54808; Initial: 7: 0.55469 | Pass
Clk Fall | ns | -- | 1 | Date (DD/MM/YY):02/22/2021 1: 0.46255; Date (DD/MM/YY):02/22/2021 2: 0.48729; Date (DD/MM/YY):02/22/2021 3: 0.50056; Date (DD/MM/YY):02/22/2021 4: 0.52573; Date (DD/MM/YY):02/22/2021 5: 0.53771; Date (DD/MM/YY):02/22/2021 6: 0.55221; Initial: 7: 0.56479 | Pass
Data Rise | ns | -- | 1 | Date (DD/MM/YY):02/22/2021 1: 0.62424; Date (DD/MM/YY):02/22/2021 2: 0.60948; Date (DD/MM/YY):02/22/2021 3: 0.59513; Date (DD/MM/YY):02/22/2021 4: 0.58088; Date (DD/MM/YY):02/22/2021 5: 0.61251; Date (DD/MM/YY):02/22/2021 6: 0.59599; Initial: 7: 0.60419 | Pass
Data Fall | ns | -- | 1 | Date (DD/MM/YY):02/22/2021 1: 0.63123; Date (DD/MM/YY):02/22/2021 2: 0.58166; Date (DD/MM/YY):02/22/2021 3: 0.62085; Date (DD/MM/YY):02/22/2021 4: 0.59061; Date (DD/MM/YY):02/22/2021 5: 0.58129; Date (DD/MM/YY):02/22/2021 6: 0.61507; Initial: 7: 0.61955 | Pass
Jitter @ IPIN = 100uA | ns | -- | 1 | Date (DD/MM/YY):02/22/2021 1: 0.17325; Date (DD/MM/YY):02/22/2021 2: 0.13413; Date (DD/MM/YY):02/22/2021 3: 0.17236; Date (DD/MM/YY):02/22/2021 4: 0.10038; Date (DD/MM/YY):02/22/2021 5: 0.06048; Date (DD/MM/YY):02/22/2021 6: 0.17205; Initial: 7: 0.06733 | Pass
Jitter @ IPIN = 1000uA | ns | -- | 1 | Date (DD/MM/YY):02/22/2021 1: 0.20227; Date (DD/MM/YY):02/22/2021 2: 0.12399; Date (DD/MM/YY):02/22/2021 3: 0.18199; Date (DD/MM/YY):02/22/2021 4: 0.07704; Date (DD/MM/YY):02/22/2021 5: 0.05844; Date (DD/MM/YY):02/22/2021 6: 0.14838; Initial: 7: 0.05374 | Pass
Pos Width max @ IPIN = 100uA |  |  |  | Date (DD/MM/YY):02/22/2021 1: 12.926; Date (DD/MM/YY):02/22/2021 2: 12.498; Date (DD/MM/YY):02/22/2021 3: 13.284; Date (DD/MM/YY):02/22/2021 4: 12.552; Date (DD/MM/YY):02/22/2021 5: 12.198; Date (DD/MM/YY):02/22/2021 6: 12.661; Initial: 7: 12.748
Pos Width max @ IPIN = 1000uA |  |  |  | Date (DD/MM/YY):02/22/2021 1: 12.985; Date (DD/MM/YY):02/22/2021 2: 12.539; Date (DD/MM/YY):02/22/2021 3: 13.322; Date (DD/MM/YY):02/22/2021 4: 12.533; Date (DD/MM/YY):02/22/2021 5: 12.203; Date (DD/MM/YY):02/22/2021 6: 12.612; Initial: 7: 12.684
Pos Width min @ IPIN = 100uA |  |  |  | Date (DD/MM/YY):02/22/2021 1: 12.247; Date (DD/MM/YY):02/22/2021 2: 11.898; Date (DD/MM/YY):02/22/2021 3: 12.57; Date (DD/MM/YY):02/22/2021 4: 12.01; Date (DD/MM/YY):02/22/2021 5: 11.795; Date (DD/MM/YY):02/22/2021 6: 11.932; Initial: 7: 12.315
Pos Width min @ IPIN = 1000uA |  |  |  | Date (DD/MM/YY):02/22/2021 1: 12.242; Date (DD/MM/YY):02/22/2021 2: 11.93; Date (DD/MM/YY):02/22/2021 3: 12.621; Date (DD/MM/YY):02/22/2021 4: 12.124; Date (DD/MM/YY):02/22/2021 5: 11.844; Date (DD/MM/YY):02/22/2021 6: 12.024; Initial: 7: 12.304
Pos Width mean @ IPIN = 100uA |  |  |  | Date (DD/MM/YY):02/22/2021 1: 12.558; Date (DD/MM/YY):02/22/2021 2: 12.2128; Date (DD/MM/YY):02/22/2021 3: 12.9324; Date (DD/MM/YY):02/22/2021 4: 12.3066; Date (DD/MM/YY):02/22/2021 5: 11.9982; Date (DD/MM/YY):02/22/2021 6: 12.2943; Initial: 7: 12.4788
Pos Width mean @ IPIN = 1000uA |  |  |  | Date (DD/MM/YY):02/22/2021 1: 12.5675; Date (DD/MM/YY):02/22/2021 2: 12.2003; Date (DD/MM/YY):02/22/2021 3: 12.9344; Date (DD/MM/YY):02/22/2021 4: 12.3189; Date (DD/MM/YY):02/22/2021 5: 11.9986; Date (DD/MM/YY):02/22/2021 6: 12.3098; Initial: 7: 12.4846
GoNoGoDuty Cycle | % | 45 | 55 | Date (DD/MM/YY):02/22/2021 1: 50.3056; Date (DD/MM/YY):02/22/2021 2: 48.7737; Date (DD/MM/YY):02/22/2021 3: 51.5785; Date (DD/MM/YY):02/22/2021 4: 49.3006; Date (DD/MM/YY):02/22/2021 5: 47.9487; Date (DD/MM/YY):02/22/2021 6: 49.2759; Initial: 7: 49.9748 | Pass
Duty Cycle @ IPIN = 100uA | % | 45 | 55 | Date (DD/MM/YY):02/22/2021 1: 50.246; Date (DD/MM/YY):02/22/2021 2: 48.891; Date (DD/MM/YY):02/22/2021 3: 51.8019; Date (DD/MM/YY):02/22/2021 4: 49.2466; Date (DD/MM/YY):02/22/2021 5: 48.0235; Date (DD/MM/YY):02/22/2021 6: 49.2183; Initial: 7: 49.9414 | Pass
Duty Cycle @ IPIN = 1000uA | % | 45 | 55 | Date (DD/MM/YY):02/22/2021 1: 50.2737; Date (DD/MM/YY):02/22/2021 2: 48.7945; Date (DD/MM/YY):02/22/2021 3: 51.7574; Date (DD/MM/YY):02/22/2021 4: 49.2794; Date (DD/MM/YY):02/22/2021 5: 48.0063; Date (DD/MM/YY):02/22/2021 6: 49.2368; Initial: 7: 49.9474 | Pass
Clock+ common mode | V | 0.9 | 1.5 | Date (DD/MM/YY):02/22/2021 1: 1.1032; Date (DD/MM/YY):02/22/2021 2: 1.0767; Date (DD/MM/YY):02/22/2021 3: 1.0789; Date (DD/MM/YY):02/22/2021 4: 1.0692; Date (DD/MM/YY):02/22/2021 5: 1.0446; Date (DD/MM/YY):02/22/2021 6: 1.035; Initial: 7: 1.0456 | Pass
Clock- common mode | V | 0.9 | 1.5 | Date (DD/MM/YY):02/22/2021 1: 1.0943; Date (DD/MM/YY):02/22/2021 2: 1.0693; Date (DD/MM/YY):02/22/2021 3: 1.0677; Date (DD/MM/YY):02/22/2021 4: 1.0566; Date (DD/MM/YY):02/22/2021 5: 1.0352; Date (DD/MM/YY):02/22/2021 6: 1.0315; Initial: 7: 1.0388 | Pass
Data+ common mode | V | 0.9 | 1.5 | Date (DD/MM/YY):02/22/2021 1: 1.1145; Date (DD/MM/YY):02/22/2021 2: 1.128; Date (DD/MM/YY):02/22/2021 3: 1.1122; Date (DD/MM/YY):02/22/2021 4: 1.0756; Date (DD/MM/YY):02/22/2021 5: 1.0843; Date (DD/MM/YY):02/22/2021 6: 1.0749; Initial: 7: 1.1027 | Pass
Data- common mode | V | 0.9 | 1.5 | Date (DD/MM/YY):02/22/2021 1: 1.0984; Date (DD/MM/YY):02/22/2021 2: 1.1017; Date (DD/MM/YY):02/22/2021 3: 1.1001; Date (DD/MM/YY):02/22/2021 4: 1.0719; Date (DD/MM/YY):02/22/2021 5: 1.0673; Date (DD/MM/YY):02/22/2021 6: 1.0682; Initial: 7: 1.0903 | Pass
Clock+ amplitude | V | 0.2 | 0.5 | Date (DD/MM/YY):02/22/2021 1: 0.37124; Date (DD/MM/YY):02/22/2021 2: 0.36242; Date (DD/MM/YY):02/22/2021 3: 0.36645; Date (DD/MM/YY):02/22/2021 4: 0.36536; Date (DD/MM/YY):02/22/2021 5: 0.35001; Date (DD/MM/YY):02/22/2021 6: 0.34174; Initial: 7: 0.35 | Pass
Clock- amplitude | V | 0.2 | 0.5 | Date (DD/MM/YY):02/22/2021 1: 0.35944; Date (DD/MM/YY):02/22/2021 2: 0.34273; Date (DD/MM/YY):02/22/2021 3: 0.34721; Date (DD/MM/YY):02/22/2021 4: 0.34204; Date (DD/MM/YY):02/22/2021 5: 0.33107; Date (DD/MM/YY):02/22/2021 6: 0.32561; Initial: 7: 0.33376 | Pass
Data+ amplitude | V | 0.2 | 0.5 | Date (DD/MM/YY):02/22/2021 1: 0.41496; Date (DD/MM/YY):02/22/2021 2: 0.42244; Date (DD/MM/YY):02/22/2021 3: 0.41725; Date (DD/MM/YY):02/22/2021 4: 0.39381; Date (DD/MM/YY):02/22/2021 5: 0.39526; Date (DD/MM/YY):02/22/2021 6: 0.38386; Initial: 7: 0.4061 | Pass
Data- amplitude | V | 0.2 | 0.5 | Date (DD/MM/YY):02/22/2021 1: 0.39728; Date (DD/MM/YY):02/22/2021 2: 0.39513; Date (DD/MM/YY):02/22/2021 3: 0.39716; Date (DD/MM/YY):02/22/2021 4: 0.38339; Date (DD/MM/YY):02/22/2021 5: 0.37344; Date (DD/MM/YY):02/22/2021 6: 0.37375; Initial: 7: 0.38476 | Pass
Bit errors @ IPIN = 100uA | -- | -- | 0 | Date (DD/MM/YY):02/22/2021 1: 0; Date (DD/MM/YY):02/22/2021 2: 0; Date (DD/MM/YY):02/22/2021 3: 0; Date (DD/MM/YY):02/22/2021 4: 0; Date (DD/MM/YY):02/22/2021 5: 0; Date (DD/MM/YY):02/22/2021 6: 0; Initial: 7: 0 | Pass
Clk errors @ IPIN = 100uA | -- | -- | 0 | Date (DD/MM/YY):02/22/2021 1: 0; Date (DD/MM/YY):02/22/2021 2: 0; Date (DD/MM/YY):02/22/2021 3: 0; Date (DD/MM/YY):02/22/2021 4: 0; Date (DD/MM/YY):02/22/2021 5: 0; Date (DD/MM/YY):02/22/2021 6: 0; Initial: 7: 0 | Pass
Bit errors @ IPIN = 1000uA | -- | -- | 0 | Date (DD/MM/YY):02/22/2021 1: 0; Date (DD/MM/YY):02/22/2021 2: 0; Date (DD/MM/YY):02/22/2021 3: 0; Date (DD/MM/YY):02/22/2021 4: 0; Date (DD/MM/YY):02/22/2021 5: 0; Date (DD/MM/YY):02/22/2021 6: 0; Initial: 7: 0 | Pass
Clk errors @ IPIN = 1000uA | -- | -- | 0 | Date (DD/MM/YY):02/22/2021 1: 0; Date (DD/MM/YY):02/22/2021 2: 0; Date (DD/MM/YY):02/22/2021 3: 0; Date (DD/MM/YY):02/22/2021 4: 0; Date (DD/MM/YY):02/22/2021 5: 0; Date (DD/MM/YY):02/22/2021 6: 0; Initial: 7: 0 | Pass
VDC Bit errors | -- | -- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Pass