## Sheet1
Opto-board Assembly Status
0
Opto-board S/N: |  |  |  | Site(check one): |  | OSU |  | Seigen
Opto-board flavor (check one): |  |  | B-Layer |  | Left |  | Right
VDC S/N | 0
0
DORIC S/N |  |  | -- | --
Step |  |  |  | Date | Initial | Comments
1) Reinforce 80-pin Connector
2) Mechanical tolerance test
3) short test
4) VDC mounting
5) DORIC mounting
6) Measure VCSEL power
7) Measure PIN respnsivity
8) Glue VCSEL opto-pack (?)
9) Glue Housing
10.1) Glue 2nd VCSEL opto-pack(?)
10.2) Glue housing (?)
11.1) Glue PIN opto-pack
11.2) Glue housing
12) Solder VCSEL opto-pack
13) Solder 2nd VCSEL opto-pack (?)
14) Solder PIN opto-pack
15) Wire bonding
16) Basic Functionality
17) Encapsulation

## Sheet2
Opto-Board QA data Sheet: Optical Power and Responsivity
Opto-board S/N
0
VCSEL Opto-pack S/N:
Power (uW) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Date | Initial | P> 500 uW & dP ,20% ?
Pre Opto-pack Burn-in |  | 2052.4504 | 1958.253 | 2038.2307 | 1991.9851 | 2022.1501 | 2099.5841 | 2029.4429 | -- | -- | --
Post Opto-pack Burn-in |  | 1923.9926 | 1847.1563 | 1905.9967 | 1882.7925 | 1893.6778 | 1972.5702 | 1910.8128 | -- | -- | --
Change (%) | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -- | -- | --
Pre Opto-board Burn-in |  | 1249.4218 | 1066.9977 | 1232.377 | 1152.3108 | 1302.841 | 1202.0732 | 1528.3867 | -- | -- | --
Change (%) |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -- | -- | --
Post Opto-board Burn-in |  | 1247.8259 | 1114.9799 | 1428.5947 | 1228.9865 | 1328.1326 | 1212.7227 | 1534.4794 | -- | -- | --
Change (%) |  | -0.12773 | 4.4969 | 15.9219 | 6.6541 | 1.9413 | 0.88592 | 0.39864 | -- | -- | --
Opto-board at 0C |  | 1457.3111 | 1272.3791 | 1591.6862 | 1335.6782 | 1480.1818 | 1329.2425 | 1631.9014 | -- | -- | --
Change (%) |  | 16.6388 | 19.2485 | 29.1558 | 15.913 | 13.6119 | 10.5792 | 6.7728 | -- | -- | --
Opto-board at 50C |  | 1269.7329 | 1085.1075 | 1393.7886 | 1172.9697 | 1292.4012 | 1149.0282 | 1464.0849 | -- | -- | --
Change (%) |  | 1.6256 | 1.6973 | 13.0976 | 1.7928 | -0.80131 | -4.4128 | -4.2072 | -- | -- | --
Opto-board QA (10C) |  | 1466.4318 | 1247.4853 | 1589.4064 | 1316.1485 | 1443.6958 | 1320.0253 | 1604.4739 | -- | -- | --
Change (%) |  | 17.3688 | 16.9155 | 28.9708 | 14.2182 | 10.8114 | 9.8124 | 4.9783 | -- | -- | --
Reception Test |  | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | -- | -- | --
VCSEL Opto-pack S/N:
Power (uW) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Date | Initial | P> 500 uW & dP ,20% ?
Pre Opto-pack Burn-in |  | 2125.4727 | 2040.2889 | 2115.8246 | 2009.1102 | 2050.1745 | 2067.8212 | 1969.6296 | -- | -- | --
Post Opto-pack Burn-in |  | 1982.8975 | 1889.0493 | 1943.386 | 1866.3421 | 1878.6837 | 1944.444 | 1825.355 | -- | -- | --
Change (%) | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -- | -- | --
Pre Opto-board Burn-in |  | 1137.5402 | 1191.0779 | 1316.1776 | 1175.2633 | 1177.6658 | 1400.3783 | 1408.0181 | -- | -- | --
Change (%) |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | -- | -- | --
Post Opto-board Burn-in |  | 1369.5859 | 1321.3045 | 1368.2875 | 1273.594 | 1275.4097 | 1413.1013 | 1358.5056 | -- | -- | --
Change (%) |  | 20.3989 | 10.9335 | 3.9592 | 8.3667 | 8.2998 | 0.90854 | -3.5165 | -- | -- | --
Opto-board at 0C |  | 1407.1696 | 1419.9708 | 1563.8592 | 1507.0307 | 1399.4459 | 1498.7537 | 1507.8907 | -- | -- | --
Change (%) |  | 23.7028 | 19.2173 | 18.8183 | 28.2292 | 18.8322 | 7.0249 | 7.0931 | -- | -- | --
Opto-board at 50C |  | 1242.6845 | 1241.0008 | 1347.6412 | 1324.7984 | 1209.9857 | 1280.7918 | 1336.0898 | -- | -- | --
Change (%) |  | 9.2431 | 4.1914 | 2.3905 | 12.7235 | 2.7444 | -8.5396 | -5.1085 | -- | -- | --
Opto-board QA (10C) |  | 1362.4117 | 1438.8605 | 1574.381 | 1469.6493 | 1361.193 | 1529.655 | 1466.0354 | -- | -- | --
Change (%) |  | 19.7682 | 20.8032 | 19.6177 | 25.0485 | 15.584 | 9.2316 | 4.1205 | -- | -- | --
Reception Test |  | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | 2000 | -- | -- | --

## Sheet3
PIN Opto-pack S/N:
Responsivity (A/W) | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | Date | Initial | R> 0.5 A/W & dR <5%
Opto-pack QA | [8 empty] | -- | -- | --
Pre-mount | [8 empty] | -- | -- | --
Change (%)

## Sheet4
Opto-board QA Data Sheet
Opto-board S/N:
0
Current consumption:
| Unit | Min | Max | I |  |  | dI <5% | Date |  | Initial |  | Pass/Fail (?)
Pre-QA | mA | 322 | 384 | 365.6006 |  |  |  | 2021-03-17 |  |  |  | Pass
QA | mA | 304 | 366 | 340.5151 |  |  |  | 2021-03-17 |  |  |  | Pass
*The allowable range for current consumption for a B-layer opto-board is 330-400 mA.
Dark current:
| Unit | Max | I |  | Date |  | Initial |  | Pass/Fail (?)
Pre-QA | uA | 1 | 0.81825 |  | 2021-03-17 |  |  |  | Pass
QA | uA | 1 | 0.55074 |  | 2021-03-17 |  |  |  | Pass
Optical
| Unit | Max | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | Pass/Fail (?)
Rise | ns | 1 | 0.46494 | 0.57375 | 0.62778 | 0.53501 | 0.69516 | 0.56089 | 0.2752 | 0.41005 | 0.66653 | 0.52585 | 0.42306 | 0.62916 | 0.40825 | 0.33204 |  |  | Pass
Fall | ns | 1 | 0.36401 | 0.34928 | 0.56909 | 0.52299 | 0.59941 | 0.59721 | 0.69278 | 0.55295 | 0.56148 | 0.45849 | 0.55603 | 0.46026 | 0.48469 | 0.57869 |  |  | Pass

## optical calc
Post Mounting @ room Temp
Power
Power 2
Power 3
channel# | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
avg | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
-10C
Power
Power 2
Power 3
channel# | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
avg | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
10C
Power
Power 2
Power 3
channel# | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
avg | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
50C
Power
Power 2
Power 3
channel# | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
| 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000
avg | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000 | 1000

## Sheet5
Electrical: |  |  |  | Date (DD/MM/YY):03/17/2021 | [5 empty] | Initial:
| Unit | Min | Max | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | Pass/fail(?)
Clk Rise | ns | -- | 1 | Date (DD/MM/YY):03/17/2021 1: 0.49328; Date (DD/MM/YY):03/17/2021 2: 0.50259; Date (DD/MM/YY):03/17/2021 3: 0.52883; Date (DD/MM/YY):03/17/2021 4: 0.52909; Date (DD/MM/YY):03/17/2021 5: 0.54579; Date (DD/MM/YY):03/17/2021 6: 0.55603; Initial: 7: 0.56714 | Pass
Clk Fall | ns | -- | 1 | Date (DD/MM/YY):03/17/2021 1: 0.48243; Date (DD/MM/YY):03/17/2021 2: 0.50043; Date (DD/MM/YY):03/17/2021 3: 0.52234; Date (DD/MM/YY):03/17/2021 4: 0.52348; Date (DD/MM/YY):03/17/2021 5: 0.544; Date (DD/MM/YY):03/17/2021 6: 0.5542; Initial: 7: 0.57095 | Pass
Data Rise | ns | -- | 1 | Date (DD/MM/YY):03/17/2021 1: 0.63683; Date (DD/MM/YY):03/17/2021 2: 0.63895; Date (DD/MM/YY):03/17/2021 3: 0.60743; Date (DD/MM/YY):03/17/2021 4: 0.58086; Date (DD/MM/YY):03/17/2021 5: 0.62738; Date (DD/MM/YY):03/17/2021 6: 0.60782; Initial: 7: 0.60273 | Pass
Data Fall | ns | -- | 1 | Date (DD/MM/YY):03/17/2021 1: 0.62759; Date (DD/MM/YY):03/17/2021 2: 0.58738; Date (DD/MM/YY):03/17/2021 3: 0.62007; Date (DD/MM/YY):03/17/2021 4: 0.59777; Date (DD/MM/YY):03/17/2021 5: 0.56305; Date (DD/MM/YY):03/17/2021 6: 0.60645; Initial: 7: 0.63005 | Pass
Jitter @ IPIN = 100uA | ns | -- | 1 | Date (DD/MM/YY):03/17/2021 1: 0.23514; Date (DD/MM/YY):03/17/2021 2: 0.07005; Date (DD/MM/YY):03/17/2021 3: 0.10888; Date (DD/MM/YY):03/17/2021 4: 0.18303; Date (DD/MM/YY):03/17/2021 5: 0.08873; Date (DD/MM/YY):03/17/2021 6: 0.10314; Initial: 7: 0.13401 | Pass
Jitter @ IPIN = 1000uA | ns | -- | 1 | Date (DD/MM/YY):03/17/2021 1: 0.2476; Date (DD/MM/YY):03/17/2021 2: 0.06242; Date (DD/MM/YY):03/17/2021 3: 0.10119; Date (DD/MM/YY):03/17/2021 4: 0.1651; Date (DD/MM/YY):03/17/2021 5: 0.05919; Date (DD/MM/YY):03/17/2021 6: 0.09789; Initial: 7: 0.13183 | Pass
Pos Width max @ IPIN = 100uA |  |  |  | Date (DD/MM/YY):03/17/2021 1: 12.698; Date (DD/MM/YY):03/17/2021 2: 12.608; Date (DD/MM/YY):03/17/2021 3: 12.395; Date (DD/MM/YY):03/17/2021 4: 12.796; Date (DD/MM/YY):03/17/2021 5: 12.755; Date (DD/MM/YY):03/17/2021 6: 12.251; Initial: 7: 12.334
Pos Width max @ IPIN = 1000uA |  |  |  | Date (DD/MM/YY):03/17/2021 1: 12.753; Date (DD/MM/YY):03/17/2021 2: 12.533; Date (DD/MM/YY):03/17/2021 3: 12.436; Date (DD/MM/YY):03/17/2021 4: 12.723; Date (DD/MM/YY):03/17/2021 5: 12.632; Date (DD/MM/YY):03/17/2021 6: 12.223; Initial: 7: 12.298
Pos Width min @ IPIN = 100uA |  |  |  | Date (DD/MM/YY):03/17/2021 1: 11.973; Date (DD/MM/YY):03/17/2021 2: 12.131; Date (DD/MM/YY):03/17/2021 3: 11.884; Date (DD/MM/YY):03/17/2021 4: 12.111; Date (DD/MM/YY):03/17/2021 5: 12.273; Date (DD/MM/YY):03/17/2021 6: 11.778; Initial: 7: 11.782
Pos Width min @ IPIN = 1000uA |  |  |  | Date (DD/MM/YY):03/17/2021 1: 11.96; Date (DD/MM/YY):03/17/2021 2: 12.138; Date (DD/MM/YY):03/17/2021 3: 11.853; Date (DD/MM/YY):03/17/2021 4: 12.189; Date (DD/MM/YY):03/17/2021 5: 12.273; Date (DD/MM/YY):03/17/2021 6: 11.78; Initial: 7: 11.772
Pos Width mean @ IPIN = 100uA |  |  |  | Date (DD/MM/YY):03/17/2021 1: 12.3248; Date (DD/MM/YY):03/17/2021 2: 12.3256; Date (DD/MM/YY):03/17/2021 3: 12.1176; Date (DD/MM/YY):03/17/2021 4: 12.4596; Date (DD/MM/YY):03/17/2021 5: 12.4802; Date (DD/MM/YY):03/17/2021 6: 12.027; Initial: 7: 12.0458
Pos Width mean @ IPIN = 1000uA |  |  |  | Date (DD/MM/YY):03/17/2021 1: 12.3328; Date (DD/MM/YY):03/17/2021 2: 12.322; Date (DD/MM/YY):03/17/2021 3: 12.1067; Date (DD/MM/YY):03/17/2021 4: 12.4539; Date (DD/MM/YY):03/17/2021 5: 12.4697; Date (DD/MM/YY):03/17/2021 6: 12.022; Initial: 7: 12.0383
GoNoGoDuty Cycle | % | 45 | 55 | Date (DD/MM/YY):03/17/2021 1: 49.4441; Date (DD/MM/YY):03/17/2021 2: 49.4699; Date (DD/MM/YY):03/17/2021 3: 48.5166; Date (DD/MM/YY):03/17/2021 4: 49.9149; Date (DD/MM/YY):03/17/2021 5: 50.0851; Date (DD/MM/YY):03/17/2021 6: 48.2563; Initial: 7: 48.2607 | Pass
Duty Cycle @ IPIN = 100uA | % | 45 | 55 | Date (DD/MM/YY):03/17/2021 1: 49.3593; Date (DD/MM/YY):03/17/2021 2: 49.3429; Date (DD/MM/YY):03/17/2021 3: 48.4993; Date (DD/MM/YY):03/17/2021 4: 49.8541; Date (DD/MM/YY):03/17/2021 5: 49.9733; Date (DD/MM/YY):03/17/2021 6: 48.143; Initial: 7: 48.2365 | Pass
Duty Cycle @ IPIN = 1000uA | % | 45 | 55 | Date (DD/MM/YY):03/17/2021 1: 49.3556; Date (DD/MM/YY):03/17/2021 2: 49.2989; Date (DD/MM/YY):03/17/2021 3: 48.4532; Date (DD/MM/YY):03/17/2021 4: 49.8127; Date (DD/MM/YY):03/17/2021 5: 49.9033; Date (DD/MM/YY):03/17/2021 6: 48.086; Initial: 7: 48.1596 | Pass
Clock+ common mode | V | 0.9 | 1.5 | Date (DD/MM/YY):03/17/2021 1: 1.027; Date (DD/MM/YY):03/17/2021 2: 1.0126; Date (DD/MM/YY):03/17/2021 3: 1.0379; Date (DD/MM/YY):03/17/2021 4: 1.0424; Date (DD/MM/YY):03/17/2021 5: 1.0402; Date (DD/MM/YY):03/17/2021 6: 1.0393; Initial: 7: 1.05 | Pass
Clock- common mode | V | 0.9 | 1.5 | Date (DD/MM/YY):03/17/2021 1: 1.0243; Date (DD/MM/YY):03/17/2021 2: 1.0086; Date (DD/MM/YY):03/17/2021 3: 1.0321; Date (DD/MM/YY):03/17/2021 4: 1.0384; Date (DD/MM/YY):03/17/2021 5: 1.0377; Date (DD/MM/YY):03/17/2021 6: 1.036; Initial: 7: 1.0458 | Pass
Data+ common mode | V | 0.9 | 1.5 | Date (DD/MM/YY):03/17/2021 1: 1.0548; Date (DD/MM/YY):03/17/2021 2: 1.0735; Date (DD/MM/YY):03/17/2021 3: 1.0795; Date (DD/MM/YY):03/17/2021 4: 1.0611; Date (DD/MM/YY):03/17/2021 5: 1.0859; Date (DD/MM/YY):03/17/2021 6: 1.0627; Initial: 7: 1.0816 | Pass
Data- common mode | V | 0.9 | 1.5 | Date (DD/MM/YY):03/17/2021 1: 1.0459; Date (DD/MM/YY):03/17/2021 2: 1.0524; Date (DD/MM/YY):03/17/2021 3: 1.0704; Date (DD/MM/YY):03/17/2021 4: 1.063; Date (DD/MM/YY):03/17/2021 5: 1.0696; Date (DD/MM/YY):03/17/2021 6: 1.0577; Initial: 7: 1.0724 | Pass
Clock+ amplitude | V | 0.2 | 0.5 | Date (DD/MM/YY):03/17/2021 1: 0.34448; Date (DD/MM/YY):03/17/2021 2: 0.33239; Date (DD/MM/YY):03/17/2021 3: 0.34585; Date (DD/MM/YY):03/17/2021 4: 0.34055; Date (DD/MM/YY):03/17/2021 5: 0.33683; Date (DD/MM/YY):03/17/2021 6: 0.33709; Initial: 7: 0.34729 | Pass
Clock- amplitude | V | 0.2 | 0.5 | Date (DD/MM/YY):03/17/2021 1: 0.31821; Date (DD/MM/YY):03/17/2021 2: 0.31014; Date (DD/MM/YY):03/17/2021 3: 0.32161; Date (DD/MM/YY):03/17/2021 4: 0.32116; Date (DD/MM/YY):03/17/2021 5: 0.32266; Date (DD/MM/YY):03/17/2021 6: 0.32301; Initial: 7: 0.33064 | Pass
Data+ amplitude | V | 0.2 | 0.5 | Date (DD/MM/YY):03/17/2021 1: 0.37951; Date (DD/MM/YY):03/17/2021 2: 0.39122; Date (DD/MM/YY):03/17/2021 3: 0.39877; Date (DD/MM/YY):03/17/2021 4: 0.36855; Date (DD/MM/YY):03/17/2021 5: 0.38948; Date (DD/MM/YY):03/17/2021 6: 0.37359; Initial: 7: 0.38758 | Pass
Data- amplitude | V | 0.2 | 0.5 | Date (DD/MM/YY):03/17/2021 1: 0.36831; Date (DD/MM/YY):03/17/2021 2: 0.36282; Date (DD/MM/YY):03/17/2021 3: 0.37842; Date (DD/MM/YY):03/17/2021 4: 0.37073; Date (DD/MM/YY):03/17/2021 5: 0.36818; Date (DD/MM/YY):03/17/2021 6: 0.36297; Initial: 7: 0.36953 | Pass
Bit errors @ IPIN = 100uA | -- | -- | 0 | Date (DD/MM/YY):03/17/2021 1: 0; Date (DD/MM/YY):03/17/2021 2: 0; Date (DD/MM/YY):03/17/2021 3: 0; Date (DD/MM/YY):03/17/2021 4: 0; Date (DD/MM/YY):03/17/2021 5: 0; Date (DD/MM/YY):03/17/2021 6: 0; Initial: 7: 0 | Pass
Clk errors @ IPIN = 100uA | -- | -- | 0 | Date (DD/MM/YY):03/17/2021 1: 0; Date (DD/MM/YY):03/17/2021 2: 0; Date (DD/MM/YY):03/17/2021 3: 0; Date (DD/MM/YY):03/17/2021 4: 0; Date (DD/MM/YY):03/17/2021 5: 0; Date (DD/MM/YY):03/17/2021 6: 0; Initial: 7: 0 | Pass
Bit errors @ IPIN = 1000uA | -- | -- | 0 | Date (DD/MM/YY):03/17/2021 1: 0; Date (DD/MM/YY):03/17/2021 2: 0; Date (DD/MM/YY):03/17/2021 3: 0; Date (DD/MM/YY):03/17/2021 4: 0; Date (DD/MM/YY):03/17/2021 5: 0; Date (DD/MM/YY):03/17/2021 6: 0; Initial: 7: 0 | Pass
Clk errors @ IPIN = 1000uA | -- | -- | 0 | Date (DD/MM/YY):03/17/2021 1: 0; Date (DD/MM/YY):03/17/2021 2: 0; Date (DD/MM/YY):03/17/2021 3: 0; Date (DD/MM/YY):03/17/2021 4: 0; Date (DD/MM/YY):03/17/2021 5: 0; Date (DD/MM/YY):03/17/2021 6: 0; Initial: 7: 0 | Pass
VDC Bit errors | -- | -- | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Pass